## Sheet1
certificate_no | N23070405/D3.03/C-036
device_name | Defibrillator Analyser
end_date | 12 07 2023
next_date | 12 07 2024
calibrated_for | Biomedical Metrology Section National Physical Laboratory, New Delhi, 110012 Customer Ref. No. Nil Date: 10-07-2023
description | Defibrillator Analyser Model No.: 7000DP, Sr. No.: 3819017 Make: Fluke Biomedical
env_conditions | Temperature: (25 ± 2) ℃ Relative Humidity: (50 ± 10) %
stds_used | 1. Digital Storage Oscilloscope (Model No.: TBS 2072; Sr. No.: C020293; Make: Tektronix); (-182.8 to -582.4) mV ± 1.5 mV (k=2.32) ; (215.8 to 2207) mV ± (1.5 to 2.0) mV (k=2.02 to 2.32); Time (sec): (19.994 ± 1 x 10-3) µs (k=2) 2. High Voltage Divider (Model No.: VD 15-8.3-A-LB-AL; Sr. No.: 170313, Make: Ross Engg. Corp.), (1000:1) 0.9975 ± 0.1 % (k=2.1) 3. Digital Multimeter (Model No. 8846A; Sr.No.: 3641001; Make: Fluke); (19.0009 ± 0.0003) Ω (k=2) 4. Defibrillator (Model No.: TEC5621; S. No.: 01273; Make: Nihon Kohden); (9.5 ± 0.1) J (k=2.16) (19.3 ± 0.1) J (k=2.00) (48.2 ± 0.2) J (k=2.13) (97.1 ± 0.3) J (k=2.00) (146.1 ± 0.5) J (k=2.06) (259.5 ± 0.8) J (k=2.05)
tracability | The standards used for calibration are traceable to National Standards, which realize the units of quantities according to the International System of Units (SI)
procedure | Calibration procedure as specified in Sub-Div # 3.03/ Doc3/CP #2
calibrated_by | VINOD KUMAR TANWAR
checked_by | VED VARUN AGRAWAL
incharge | Dr. RAJESH
issued_by

## Sheet2
result_table | Sr. No. | Energy applied (J) to UUC | Calculated (J) level | Measured Energy (J) level of UUC | Energy correction (J) (Ecalculated-Emeasured) | Uncertainty (J) | Coverage Factor (k)
| 1 | 10 | 9.5 | 9.6 | (-) 0.1 | ± 0.1 | 2
| 2 | 20 | 19.3 | 19.4 | (-) 0.1 | ± 0.1 | 2
| 3 | 50 | 48.2 | 48.4 | (-) 0.2 | ± 0.2 | 2
| 4 | 100 | 97.1 | 97.4 | (-) 0.3 | ± 0.3 | 2
| 5 | 150 | 146.1 | 146.4 | (-) 0.3 | ± 0.5 | 2
| 6 | 270 | 259.5 | 260.3 | (-) 0.8 | ± 0.9 | 2
| Sr. No. | Discharge Time | Calculated Discharge Time (ms) | Measured Discharge Time (ms) | Uncertainty (J) | Coverage Factor (k)
| 1 | Discharge Time (Positive Pulse Width) | 6.6 | 6.4 | ± 0.1 | 2.17
| 2 | Discharge Time (Negative Pulse Wdith) | 4.4 | 4.5 | ± 0.1 | 2
| Sr. No. | Defribillator Analyser Load Resistance (Ω) | Expected Value (Ω) | Measured Value (Ω) | Uncertainty (Ω) | Coverage Factor (k)
| 1 | 50 | 50 | 50.0179 | ± 0.00041 | 2.05
| Sr. No. | Defribillator Analyser Load Resistance (Ω) | Expected Value (Ω) | Measured Value (Ω) | Uncertainty (Ω) | Coverage Factor (k)
| 1 | 50 | 50 | 51.0128 | ± 0.0005 | 2.11
result_desc | The expanded uncertainty in wavelength measurement at NPL. is ±0.5 m. The reported uncertainty is at Coverage factor 1-2, which corresponds to a coverage probability of approximately 95% for a normal distribution
calibration_date | 12.07.2023
remarks | (i) The measured values of peak wavelength are representation of the nearest reference Standard spectral lines as shown by the Spectrometer at the interval of each 0.1 mm. (ii) NPL identification No. of the Spectrometer for wavelength measurement is 402/0PT/202